Standard Operating Procedure

Electrical Assembly of the Automated pTLC Spotter Device

| Document Control |  |
| --- | --- |
| SOP Title | Electrical Component Assembly – pTLC Spotter Device |
| Prepared By | Joseph Thomas |
| Reviewed By | Prof. Chukwunonso Nwabufo |
| Related Documents | Electromechanical Workflow of the Diagnostic Device; Device Wiring Schematic |
| Revision | 1.0 |

# 1. Purpose

This SOP defines the step-by-step procedure for assembling and wiring the electrical and electronic components of the automated pTLC Spotter Device. It covers the controller subsystem (Raspberry Pi 4 and ESP32), motion control electronics (stepper and servo drivers), sensing components (IR temperature sensor), actuation components (linear actuators, servos, heating pads, LED), and the imaging subsystem (Raspberry Pi HQ Camera). This procedure should be followed in conjunction with the device wiring schematic and the Electromechanical Workflow reference document.

# 2. Scope

This procedure applies to technicians and engineers responsible for assembling the electrical subsystem of the pTLC Spotter Device prior to mechanical integration and functional testing. It assumes the mechanical/structural assembly (chassis, actuator mounting, servo brackets) has already been completed per the mechanical design documentation.

# 3. Safety Precautions

CAUTION: Disconnect all power sources (mains adapter and any connected battery/USB power bank) before making or modifying any wiring connections. Never wire components while the system is powered.

- Wear an anti-static wrist strap or otherwise ground yourself before handling the Raspberry Pi 4, ESP32, or any exposed circuit boards to prevent ESD damage.

- Confirm supply voltage and polarity before connecting any component. Reversed polarity can damage the Raspberry Pi, ESP32, servos, or drivers.

- Keep the workspace dry and free of conductive debris (metal shavings, loose screws) before power-up.

- Verify current ratings of wires and connectors against the stepper motor and heating pad draw before final assembly; undersized wiring can overheat.

- Do not exceed the rated voltage of the ST3020 servos (6–14 V, nominal 12 V) or the SG90 servo (nominal 5 V).

- Allow the heating pads to cool and disconnect them before touching the heating plate assembly during testing.

# 4. Materials and Tools Required

## 4.1 Electronic Components

| Component | Qty | Notes |
| --- | --- | --- |
| Raspberry Pi 4 | 1 | Main system controller |
| ESP32 development board | 1 | Low-level motion/servo control |
| NEMA 14 (3534) two-phase stepper motors | 2 | Drives Actuator 1 and Actuator 2 |
| Stepper motor drivers | 2 | One per NEMA 14 stepper |
| Waveshare ST3020 serial-bus servo | 2 | Needle servo, Stage transfer servo |
| SG90 micro servo | 1 | Dipping servo |
| Servo driver / serial bus board | 1 | Interfaces ST3020 servos to controller |
| MOSFET | 1 | Switches LED on |
| Non-contact IR temperature sensor | 1 | Monitors microwell plate temperature |
| 7 W adhesive heating pads | 1 set | Mounted to heating plate, Part 8 |
| High-intensity LED (excitation) | 1 | Fluorescence excitation source |
| Raspberry Pi HQ Camera + 6 mm lens | 1 | Imaging module, Part 11 |
| DC power supply/adapter | 1 | Sized for combined system load |
| Hookup wire, JST/Dupont connectors, heat-shrink | as required | Wiring harness |

## 4.2 Tools

- Phillips and flathead screwdriver set

- Wire strippers and crimping tool

- Soldering iron and solder (if permanent joints are required)

- Digital multimeter

- Anti-static wrist strap

- Small zip ties / cable management clips

- Heat gun (for heat-shrink tubing)

- Laptop with Raspberry Pi OS imaging tool and ESP32 IDE (e.g., Arduino IDE or PlatformIO) installed

# 5. Pre-Assembly Checks

- Confirm the mechanical/structural build (chassis, actuator brackets, servo mounts) is complete and matches the mechanical design documentation.

- Lay out all electronic components and cross-check against the Bill of Materials in Section 4.1.

- Inspect the wiring schematic to confirm the current hardware revision matches the diagram provided.

- Bench-test the Raspberry Pi 4 and ESP32 individually (power up, verify boot) before integrating them into the chassis.

- Verify the DC power supply output voltage and current rating with a multimeter before connecting to any board.

# 6. Step-by-Step Electrical Assembly

## 6.1 Mount and Position the Controllers

- Install the Raspberry Pi 4 into its designated mounting standoffs inside the bottom cover (Part 13) enclosure.

- Install the ESP32 development board on its mount adjacent to the Raspberry Pi 4, allowing clear access to its USB and GPIO headers.

- Route a short USB or UART cable between the Raspberry Pi 4 and ESP32 per the schematic, securing it with a cable clip to prevent strain.

## 6.2 Wire the Power Distribution

- Connect the main DC power input to the power distribution point identified on the schematic.

- Wire the 5V/3A regulated supply line to the Raspberry Pi 4 power input.

- Wire the 6–14 V servo supply line to the ST3020 servo driver/bus board, observing correct polarity.

- Wire the 5 V supply line to the SG90 micro servo signal/power header.

- Wire the stepper motor driver power rails to the appropriate voltage rail specified for the NEMA 14 stepper motors.

- Confirm all grounds (Raspberry Pi, ESP32, stepper drivers, servo driver, relay module) are tied to a common ground per the schematic before proceeding.

CAUTION: Do not connect the servo or stepper power rails to the Raspberry Pi's 5V GPIO pins. These loads must be powered from the dedicated regulated supply lines shown in the schematic, not the Pi's onboard regulator.

## 6.3 Wire the Stepper Motor (Linear Actuator) Circuits

- Connect Actuator 1's NEMA 14 stepper motor leads to Stepper Driver 1 per the coil pairings shown on the schematic.

- Connect Actuator 2's NEMA 14 stepper motor leads to Stepper Driver 2 in the same manner.

- Wire the STEP, DIR, and ENABLE control lines from each stepper driver to the designated ESP32 GPIO pins shown on the schematic.

- Set the current-limit potentiometer (or digital current setting) on each stepper driver to match the NEMA 14 rated current, then verify with a multimeter across the driver's reference pin.

## 6.4 Wire the Servo Circuits

- Connect the Needle Servo (ST3020) and Stage Transfer Servo (ST3020) to the serial-bus servo driver board using the servo bus cable connectors.

- Connect the servo bus driver board's data line to the designated UART/TTL pin on the ESP32 (or Raspberry Pi, per schematic).

- Connect the SG90 Dipping Servo signal wire directly to its designated PWM-capable GPIO pin; connect its power and ground leads to the 5 V rail and common ground.

NOTE: The ST3020 servos use a closed-loop serial-bus protocol (position, torque, temperature feedback) and must be addressed with the correct servo ID as configured in firmware; the SG90 is a simple 3-wire PWM servo.

## 6.5 Wire the Heating and Lighting Circuits

- Connect the 7 W adhesive heating pads (mounted on the heating plate, Part 8) to the relay module output channel designated for heating control.

- Wire the relay module's control input to the ESP32/Raspberry Pi GPIO pin identified in the schematic for temperature control.

- Connect the non-contact IR temperature sensor to its designated communication bus (e.g., I2C) on the Raspberry Pi or ESP32, including pull-up resistors if not already present on the sensor breakout.

- Connect the high-intensity excitation LED to its relay/driver output channel, and wire the relay control input to the designated GPIO pin for illumination control.

CAUTION: Route heating pad wiring away from the solvent chamber and any wiring near the imaging module to avoid heat damage or fire risk near solvent vapors.

## 6.6 Wire the Imaging Module

- Connect the Raspberry Pi HQ Camera (Part 11) to the Raspberry Pi 4 CSI camera port using the ribbon cable, ensuring the contacts face the correct orientation.

- Route the ribbon cable away from moving parts (actuators, servo arms) and secure it with a cable clip.

- If the excitation LED or emission filter assembly includes any powered components (e.g., a filter actuator), wire them per the schematic to their designated control output.

## 6.7 Cable Management and Enclosure

- Bundle and secure all wiring harnesses with zip ties, keeping motor/servo power lines separated from low-voltage signal lines where possible to reduce electrical noise.

- Label each connector or wire bundle to match its designation on the schematic (e.g., “ACT1”, “ACT2”, “NEEDLE-SERVO”, “DIP-SERVO”, “TRANSFER-SERVO”, “HEAT”, “LED”, “IR-TEMP”).

- Confirm no wires obstruct the travel path of either linear actuator or any servo arm through a manual, unpowered range-of-motion check.

- Install the bottom cover (Part 13) over the electronics once wiring is confirmed correct, then fasten the main chassis to it.

# 7. Post-Assembly Verification and Testing

- Perform a visual inspection of all connections against the schematic before applying power.

- Use a multimeter to verify continuity on ground connections and confirm no short circuits exist between power rails.

- Apply power and confirm the Raspberry Pi 4 and ESP32 boot correctly.

- Individually jog each stepper-driven actuator through a short, low-speed test move and confirm correct direction of travel.

- Individually command each servo (Needle, Stage Transfer, Dipping) through a small test motion and confirm correct direction and smooth operation.

- Verify the IR temperature sensor returns a plausible ambient temperature reading.

- Verify the relay-controlled heating pad and LED energize and de-energize correctly on command.

- Capture a test image using the Raspberry Pi HQ Camera to confirm the imaging module is functional.

- Record all verification results on the Electrical Assembly Checklist (Section 9) and route to the reviewer for sign-off.

# 8. Troubleshooting Reference

| Symptom | Likely Cause | Corrective Action |
| --- | --- | --- |
| Raspberry Pi or ESP32 fails to boot | Incorrect supply voltage/polarity; loose power connector | Re-check power wiring against schematic; verify voltage with multimeter |
| Stepper motor does not move or moves erratically | Miswired coil pairs; incorrect current-limit setting; loose STEP/DIR wiring | Verify coil pairing against motor datasheet and schematic; re-check driver current setting |
| Servo does not respond | Incorrect bus ID; power rail not connected; signal line miswired | Confirm servo ID in firmware; verify 6–14V (ST3020) or 5V (SG90) supply; re-check signal wiring |
| Heating pad does not energize | GPIO control line not connected; blown fuse | Verify relay wiring and control signal; check inline fuse if present |
| No camera image / camera not detected | Ribbon cable inserted incorrectly or damaged | Re-seat CSI ribbon cable, verify orientation, inspect for damage |
| IR temperature sensor reads incorrectly or not detected | Bus wiring error; missing pull-up resistors; address conflict | Verify I2C/bus wiring and pull-ups; confirm sensor address |

# 9. Electrical Assembly Checklist (Sign-Off)

| Item | Verified (Y/N) | Initials / Date |
| --- | --- | --- |
| All connections match wiring schematic |  |  |
| Power rails verified with multimeter prior to power-up |  |  |
| Raspberry Pi 4 and ESP32 boot successfully |  |  |
| Actuator 1 stepper motion verified |  |  |
| Actuator 2 stepper motion verified |  |  |
| Needle Servo (ST3020) motion verified |  |  |
| Stage Transfer Servo (ST3020) motion verified |  |  |
| Dipping Servo (SG90) motion verified |  |  |
| IR temperature sensor reading verified |  |  |
| Heating pad relay control verified |  |  |
| Excitation LED relay control verified |  |  |
| Camera image capture verified |  |  |
| Cable management / labeling complete |  |  |

Assembled by (Name/Signature): ______________________________     Date: ______________

Reviewed by (Name/Signature): ______________________________     Date: ______________

# 10. References

- Electromechanical Workflow of the Diagnostic Device — pTLC Spotter Device (system, mechanical, and subsystem design document).

- Device Wiring Schematic (electrical connection diagram for Raspberry Pi 4, ESP32, drivers, servos, sensors, and imaging module).